## Slide 1: DESIGN OF A SPATIALLY RESOLVED VUV SPECTROSCOPY SYSTEM FOR SHOCK TUBE FLOWS
Maïlys Buquet, Alex Glenn, Peter Collen, Benjamin Williams, Matthew McGilvray and Tobias Hermann
University of Oxford
Tuesday 13 – Sept – 2022
RHTG Conference

## Slide 2: What? Shock Layer and Radiative Heating
Fig 1: Schematic of the high-temperature regions
in an entry-body flowfield [1]
Through shock
    Temperature increase
        Ionization & dissociation
          Chemical reactions induce radiation
→ Radiation causes additional heating
[1] Hypersonic and High Temperature Gas Dynamics, John D. Anderson Jr

## Slide 3: What? Shock Layer and Radiative Heating
[2] Characterization of Stagnation-Point Heat Flux for Earth Entry, A. M. Brandis and C. O. Johnston, 45th AIAA Plasmadynamics and Lasers Conference, 2014
Fig.2: Share of radiative and convective heat fluxes at stagnation point for different velocities and nose radii. LAURA/HARA simulations, earth atmosphere. [2]
Through shock
    Temperature increase
        Ionization & dissociation
          Chemical reactions induce radiation
→ Radiation causes additional heating
At speeds > 10 km.s-1 :
Radiation becomes substantial / main source of heat load
→ non-equilibrium region
→ high energy transition in the vacuum ultra-violet (VUV) can
become dominant
(up to 85% of the total post-shock radiance for high enthalpy shock tube tests [3])
[3] Air Radiation in Superorbital Expanding Flow, Han Wei, Thesis
Large
uncertainties
→
Largest source of
uncertainties
→
Due to complexity of measuring rad. in this wavelength – fully absorbed by atmosphere

## Slide 4: What? Shock Layer and Radiative Heating
[2] Characterization of Stagnation-Point Heat Flux for Earth Entry, A. M. Brandis and C. O. Johnston, 45th AIAA Plasmadynamics and Lasers Conference, 2014
Radiation becomes substantial / main source of heat load
→ non-equilibrium region
→ high energy transition in the vacuum ultra-violet (VUV) can
become dominant
Large
uncertainties
→
Largest source of
uncertainties
→
→ Overdesign of thermal protection systems
Reduced travel distance / useful load
Increased cost / waste
Better understanding → better model → more sensible design
Obtain spectra and absolute radiance from VUV:
very valuable
→ species number densities
→ temperature
Fig.2: Share of radiative and convective heat fluxes at stagnation point for different velocities and nose radii. LAURA/HARA simulations, earth atmosphere. [2]

## Slide 5: How? Shock tube to recreate conditions
[1] Hypersonic and High Temperature Gas Dynamics, John D. Anderson Jr
Fig 1: Schematic of the high-temperature regions
in an entry-body flowfield [1]
≡
Fig 4: T6 shock tube schematic, modified from [4]
[4] High Enthalpy Effects in Transpiration Cooled System, S. D. Subiah
T6 shock tube mode
High speed, high enthalpy facility

## Slide 6: How? Spectroscopy to investigate the flow
Andor iStar sCMOS
Vacuum chambers
containing mirrors
Spectrometer
Intensified
camera
[5] Measurements of Radiating Hypervelocity Air Shock Layers in the T6 Free-piston Driven Shock Tube, P. L. Collen, L. J. Doherty, M. McGilvray
Fig 5: Example spectra (T6s52, 10.18 km.s−1) red spectrum [5]
Species
Temperature
Species number
densities
Tunnel
axis
Shock direction
Shock

## Slide 7: How? Spectroscopy set-up on T6
→ Allows to keep
system under
vacuum during
tunnel turn-around
Spectrograph
Vacuum chambers
Camera
Shock
tube
Turbopump
Test section

## Slide 8: How? Spectroscopy set-up on T6
→ Can perform calibration during turn-around using a deuterium lamp

## Slide 9: How? Optical system design
[6] Collection Optics for Imaging Spectroscopy of an Electric Arc Shock Tube, Greenberg, Reid, Cruden, Brett, Grinstead, Jay, Yeung, Dickson
Shock
Tunnel
axis
Blur
θ
Chief rays from all object points
perpendicular to shock tube
Blurring effects due to
non-telecentricity [6]

## Slide 10: How? Optical system design
Bunch of constraints
Optical
Magnification
Matching F#
Focal lengths
90deg turn for camera
Minimum number of mirrors, smallest angles possible…
Geometrical
Must fit in boxes
Must not be on optical path
Tilting mirror makes it harder to make them fit
Optimisation approach with a code
Position and geometry of mirrors
Ray-tracing simulation using OPTOMETRIKA [7]
[7] Optometrika (https://github.com/caiuspetronius/Optometrika), Yury, 2002
Tunnel axis
Spec.
slit

## Slide 11: How? Optical system characterisation
Cylindrical mirror
Helps rays lying in perpendicular planes on the optical axis to focus at different distance
Strategy used for the VUV
spectroscopy set-up in EAST [6]
[6] Collection Optics for Imaging Spectroscopy of an Electric Arc Shock Tube, Greenberg, Reid, Cruden, Brett, Grinstead, Jay, Yeung, Dickson

## Slide 12: How? Optical system characterisation
θ > 40deg
θ = 30deg
Magnification:
M = 10.5 → M = 7.2

## Slide 13: How? Calibration
→ Some effects obtained from simulations
SPACE
→
LSF
SRF

## Slide 14: How? Calibration
SPACE
WAVELENGTH –
– INTENSITY
→ Some effects obtained experimentally
Mercury lamp spectra
Deuterium lamp spectra

## Slide 15: What’s to come… hopefully
→ Calibration applied to Poshax [8] and NEQAIR [9] simulations
→ Simulations ran to select wind-tunnel conditions
→ EAST conditions - gave good VUV data
→ Fire II re-entry datapoint
→ Data point from previous T6 campaign
[8] Eilmer 4.0 flow simulation program. P. A. Jacobs and R. J. Gollan, 2022.
[9] NEQAIR - Nonequilibrium Radiative Transport and Spectra Program, NASA
P∞ = 26.66 Pa
u∞ = 10.66 km.s-1
P∞ = 2.08 Pa
u∞ = 11.36 km.s-1
P∞ = 6.66 Pa
u∞ = 11.38 km.s-1

## Slide 16: Summary
So far…
→ Mechanical system designed,
almost fully assembled
→ Optics designed and characterised
→ Couple of test conditions selected
To come…
→ Characterised rest of optics (mirrors)
→ Pump down full volume
→ T6 test campaign to obtain spectra

## Slide 17: Thank you for listening!