## Slide 1: Characterization and Calibration of the COSSTA Radiative Heating Measurements
P. zur Nieden, T. Thiele, N. Wendel, F. Siebe, D. Neeb, A. Gülhan
Institute of Aerodynamics and Flow Technology
Supersonic and Hypersonic Technologies Department

## Slide 2: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 3: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 4: What is COSSTA?
DrEAM = “Dragonfly Entry Aerosciences Measurements”
Collaboration: DLR & NASA
Instrumentation for atmospheric entry at Titan
Provide critical aerothermodynamic data for TPS
“COmbined Sensor System for Titan Atmosphere” (COSSTA)
Based on COMARS (Schiaparelli)
3 sensor clusters (DLR) containing:
  2 total heat flux sensors
  3 pressure transducers
  7 radiometers
?

## Slide 5: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 6: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 7: Titan‘s Atmosphere
N2
CH4
N2
N2

## Slide 8: Atmospheric Entry at Titan
N2
CH4
N2
N2
CN

## Slide 9: Convection vs Radiation
CN
Backshell Shoulder

## Slide 10: Convection vs Radiation
CN
Notes: Abundance of methane

## Slide 11: Atmospheric Entry at Titan
CN
Nonequilibrium CN concentrations
Accurate modeling extremely difficult!
Important to understand these mechanisms for effective TPS design
Goal:
  Enhance our understanding of Titan atmospheric entry
  Contribute valuable data to the broader field of entry science

## Slide 12: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 13: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 14: Radiometer Distribution on Backshell of Dragonfly EV
7 thermopile sensors:
2 Narrowband radiometers
CN red / CN violet
1 Broadband radiometer
CN radiation
2 Narrowband radiometers
CN red / CN violet
2 Narrowband radiometers
CN red / CN violet
Notes: Sapphire window

## Slide 15: Radiative Heat Flux
spectral radiance
Hemispherical heat flux:
Short / long pass filter .
Thermopile sensor . with Sapphire filter .
Radiometer body .
Notes: Particles from pyrolysis / ablation

## Slide 16: Interpretation of Measured Heat Flux 1
Irradiance for specific
angle of aperture
filters
spectral radiance
Hemispherical heat flux:

## Slide 17: Angular Characterization
Blackbody Landcal R1500T
1100 °C
Rotation stage
Thermopile sensor
Irradiance for specific
angle of aperture
filters

## Slide 18: Angular Characterization
Irradiance for specific
angle of aperture
filters

## Slide 19: Interpretation of Measured Heat Flux 2
Measure of radiance
Irradiance for specific
angle of aperture
filters

## Slide 20: Interpretation of Measured Heat Flux (Recap)
Radiance
Irradiance
Hemispherical heat flux

## Slide 21: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 22: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 23: Angular Characterization: Test Setup
200 cm
Rotation stage
Radiometer body
with thermopile
Blackbody Landcal R1500T
1100 °C
Rotation stage

## Slide 24: Test Parameters
Surface properties
Length
Angle of aperture

## Slide 25: Surface Properties
Copper (smooth/rough)
Titanium (smooth/rough)
Notes: Abrasive blasting
Defined sandblast
Reduce reflections in the well

## Slide 26: Length
Short ~ 8 mm
Long ~ 40 mm

## Slide 27: Angle of Aperture
40°
10°

## Slide 28: Results: Surface Properties

## Slide 29: Results: Surface Properties

## Slide 30: Results: Length

## Slide 31: Results: Length

## Slide 32: Results: Length

## Slide 33: Results: Angle of Aperture

## Slide 34: Results: Angle of Aperture

## Slide 35: Results: Angle of Aperture

## Slide 36: Results: Angle of Aperture

## Slide 37: Additional Results: Temperature

## Slide 38: Conclusion for Heat Flux Measurement
Influence of material (copper vs titanium) is not significant
Great importance of surface roughness
A shorter distance of the sensor from the capsule surface yields a preferable (plateau-shaped) angular characteristic suitable for measurement of irradiance (rather than radiance)
A greater angle of aperture (40° vs 10°) yields a preferable (plateau-shaped) angular characteristic suitable for measurement of irradiance (rather than radiance)
This is particularly true for a long distance of the sensor from the capsule surface. A shorter distance mitigates this effect.

## Slide 39: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 40: Overview
What is COSSTA?
Atmospheric Entry at Titan
Radiative Heating Measurements
Radiometer Characterization
  Set-up
  Parameters
  Results
Outlook

## Slide 41: Outlook
Angular characterization of actual flight sensors mounted in COSSTA sensor bodies
  Blackbody cavity radiator for broadband radiometers
  Measured angular characteristic of narrowband radiometers depends on light source – blackbody vs. red/UV LED
    Presumably due to wavelength-dependent reflection in the well
    Possible to use LEDs
Calibration of the absolute irradiance for all radiometers
  Blackbody surface radiator
  Up to 40 W/cm²

## Slide 42: Outlook (continued)
Test thermopile sensors for resistance at low temperatures of down to -155 °C
  Cryogenic thermal vacuum chamber
  Performance check of thermopile channels after exposure to low temperatures
  Temperature measurement during temperature cycles
Calibration of the wavelength sensitivity of the thermopile sensors
Dragonfly launch scheduled for 2028
  Enhance our understanding of Titan atmospheric entry
  Contribute valuable data to the broader field of entry science

## Slide 43: Questions…?